## Slide 1: ESS medium beta cavity recovery
L. Monaco
INFN Milano – LASA
for the SRF group

## Slide 2: INFN-LASA Medium Beta Section Contribution
LASA responsible for this Italian in-kind contribution to ESS
36 (+2) 704.42 MHz MB cavities (b = 0.67), fully treated and prepared for VT at vendor, ready for installation in cryomodule
  “built to print” strategy
  R&D on prototypes necessary for setup treatment at vendor and QA/QC control
  PED art. 4.3 included from beginning
Five entities (INFN, DESY, CEA, ESS, ZRI) involved in the process
Well defined interfaces are key points for success of collaboration
QA/QC is mandatory to supervise the whole process

## Slide 3: INFN-LASA MB Production Cycle

## Slide 4: INFN-LASA MB Production Cycle & QA/QC

## Slide 5: Nb Procurement and ECS QC
EDDY scanning results
About 4 tons of Nb sheets delivered in 3 batches and inspected at the company before delivery
Eddy Current Scanning of all 456 sheets at DESY (no foreign materials found)
Niobium ingot and forged piece at Ningxia OTIC
Sheet inspection at Ningxia OTIC
Nb grain size, RRR and mechanical properties measurement
Eddy current scanning at DESY
SEM after final polishing
EDDY scanning 24 spare sheets

## Slide 6: Actual delivery status
28 cavities delivered within specifications
  all the cavities overcome specifications in term of Q0 by at least a factor 2.
  all cavities have negligible field emission
3 x ESS Goal
ESS Goal
1.5 x ESS Goal

## Slide 7: HPR for FE suppression: M006 qualification
  HPR cure FE (only 2 cavities affected on the overall 36)
  M006 VT:
    1st VT: naked (after local grinding) -> strong MP not conditioned and high FE
    2nd VT: series VT (with tank) -> strong MP not conditioned and high FE
    3rd VT: series VT after optimized HPR -> no MP and no FE
HPR retreatment

## Slide 8: Bulk BCP new cycle: BCP grooves on MC tube
  Grooves on the MC tubes after Bulk BCP (done in two steps)
  Reason: BCP acid stagnant and possible air bubble when the cavity is in vertical position with the MC on the top
  How we recovered:
    Cavities already treated with Bulk BCP:
      Local grinding at the MC tube and 50mm BCP
    Cavities still to be treated with Bulk BCP:
      Reversing of the two main steps (starting with MC up)
      if grooves visible -> grinding and then 2nd BCP step
2nd (90mm)
1st (110mm)
Before and after grinding
MC
MC

## Slide 9: Actual delivery status
28 cavities delivered within specifications
8 (+2) cavities are under qualification process
Cavity vendor (Ettore Zanon) financial crisis results in its failure and the acquisition of the company by SIMIC that created the Zanon Research and Innovation keeping the know-how
This, together with COVID’19, influenced the quality of the cavities resulting in 10 cavities not yet qualified

## Slide 10: Batch of available cavities for qualification
7 cavities jacketed
  M015
  M021
  M022
  M024
  M026
  M027
  M028
3 naked cavities
  M017
  M024
  M037
No FE detected!

## Slide 11: Production «review»
  Buffered Chemical Polishing (BCP) on MB Cavities is critical due to the large cell radius and steep wall angle
  We have observed grooves, deep and with high spatial frequency, on the cavity walls in some of the cavities not yet qualified. These might be the cause of not proper final cleaning or a source of dissipation
  No field emission has been observed on these cavities
  The new treatment approaches aims to improve final surface finishing to improve smoothness and/or cleaness.
Grooves pattern (from replica)

## Slide 12: Boundary on additional cavity treatment
Jacketed cavities
  Limited frequency span
    Tank constrains cavity -> material removal limited to lower frequency boundary
  Field Flatness conservation
    Material removal not perfectly uniform -> possible FF change w/o tuning possibility
Naked cavities
  Final cavity length
    Further Nb material removal -> longer cavity
  Cavity sensitivities (LFD, pressure, etc.)
    Further Nb material removal -> thinner cavity wall

## Slide 13: Naked cavities

## Slide 14: New treatment approach
  In collaboration with Zanon Research & Innovation, we have agreed to treat an ESS MB cavity with Electro Polishing process, alternative to BCP used for series production.
  ZRI has long experience on EP starting from European XFEL, LCLS-II and, recently, with a cavity of the LB type for the PIP-II project.
  The EP process works differently w.r.t. BCP for material removal, acting more efficiently on edges than on flat surfaces. For this reason, we applied this treatment aiming to smooth the grooves present in our cavity. As a consequence of the increase smoothness, also the cleanness of the cavity should improve.

## Slide 15: M024
  M024 has been selected as “test” cavity for EP treatment.
  “Series” BCP treatment
    Bulk: 228.7 mm
    Grinding MC tube: 50 mm
    Final BCP: 20 mm
    XRF: no result
  New EP treatment
    Bulk EP: 104.3 mm
    Heat treatment: 600 °C/10h
    Final EP: 10.3 mm
…. but ….
BCP
EP

## Slide 16: M024
Measurement after VT test
  FF: 80 %
  Mechanics: out of tollerance
    Dl = -0.89 mm length
Cavity was set back into tollerance and proceeded to integration, informing INFN only at integration already started.
  Detuned cell was the limiting one.
ESS Field Flatness specification
-14 %

## Slide 17: M024
Final «BCP» treatment
  BCP: 10 mm
Eacc = 13.7 MV/m
Q0 = 1.4e10
 - 15%
Before
After

## Slide 18: M024
Lessons learnt
  EP process successful in improving significantly cavity performance
    Maximum accelerating field improved by 80 %
  New process flow for EP introduces unexpected deformation to the cavity
    Activity is on-going to identify possible sources
    Additional control and approval actions introduced in the cavity life-cycle to handle and prevent them
  Unstable etching rate between treatments
EP effect for cavity recovery (1st vs 3rd VT)

## Slide 19: M017
Given the initial successful results of M024, the same treatment was applied to cavity M017 and M037
  “Series” BCP treatment
    Bulk: 210.0 mm
    Grinding MC tube: 50 mm
    Grinding defect: 50 mm
    Final BCP: 20 mm
  New EP treatment
    Bulk EP: 109.7 mm
    Heat treatment: 600 °C/10h
    Final EP: 10.9 mm
BCP
EP

## Slide 20: M017
Measurement after VT test
  FF: 92.2 %
  Mechanics: out of tollerance
    Dl = -0.61 mm length
Cavity is on hold!
Eacc = 11.5 MV/m
Q0 = 1.6e10

## Slide 21: M037
  “Series” BCP treatment
    Bulk: 206.2 mm
    Grinding defects: 50 mm
    Final BCP: 20 mm
  Modified BCP (High Flow)
    Bulk BCP: 62.4 mm
    Final BCP: 20 mm
  New EP treatment
    Bulk EP: 109.7 mm
    Heat treatment: 600 °C/10h
    Final EP: 10.0 mm
BCP
BCP High Flow
BCP High Flow
EP

## Slide 22: M037
Measurement after VT test
  FF: 89.4 %
  Mechanics: OK
    Dl = -0.6 mm length
Cavity is on hold!
Eacc = 15.7 MV/m
Q0 = 1.5e10
ESS Field Flatness specification
-5 %

## Slide 23: Jacketed cavities

## Slide 24: New treatment approach
  In collaboration with Argonne National Laboratory, we have agreed to treat an ESS MB cavity with Rotational BCP process, alternative to BCP used for series production.
  Jacketed cavities are already integrated in the He-Tank and this limits the possibility for frequency variations before a removal of the He-Tank is necessary.
  We have chosen horizontal rotational BCP as a possible treatment for recovery similar cavities being alternative to our original one and given the good results obtained on similar cavities.

## Slide 25: M028
Cavity was treated at ANL removing 25 mm.
ZRI prepared the cavity for Vertical Test
  Based on ANL experience, Viton sealings were used on the cavity to protect the inner surface during transportation instead of Hex Aluminum.
  Flanges were carefully cleaned at Zanon before preparation for VT
  Only HPR, no final “BCP” not to change surface finishing
Cavity was tested at CEA
  Test interrupted for huge field emission
Cavity is now at ZRI for new cleaning process (US + soap, HPR) before new VT

## Slide 26: M028
“Series” BCP treatment
  Bulk: 206.2 mm
  Grinding defects: 50 mm
  Final BCP: 20 mm
Additional BCP (FE in previous VT test)
  Final BCP: 20 mm
  Reduced maximum gradient
New HPR
  No changes
Rotational BCP
  25 mm
  HPR
  Test at CEA
    Huge FE
BCP
20 mm BCP
20 mm BCP
Rotational BCP

## Slide 27: New cavities production

## Slide 28: New treatment process - ElectroPolishing
Four new cavities are being produced at ZRI with the new EP process flow.
All cavities have been mechanical fabricated and treated with bulk EP process.
We originally plan to have two of them tested naked to qualify the new EP treatment

## Slide 29: M039
EP treatment
  Grinding at EQ after EBW
  Bulk: 160 mm
  Final EP: 10 mm
Test naked
  EaccMax = 22.2 MV/m
  Q0Max = 4.4e9
  Q0 EaccGoal = 1.3e10
Cavity on hold before Integration
  Waiting FF and mechanical measurements

## Slide 30: M040
EP treatment
  Bulk: 160 mm
No significant defects
Few “lines” in the equatorial region
Replicas do not help in understanding
We plan for a VT test naked middle of October (wk. 42)
After EBW
After EP

## Slide 31: M042
EP treatment
  Bulk: 159.9 mm
Defect after EBW grinded and not visible
Defect re-appears after “bulk” EP
After EBW
After EP
After EBW + grinding

## Slide 32: M042
3-D Microscopy
  Defects Dimensions
  Morphology investigation
    Note different finishing of the front surface
Transverse section profile
Border section profile
100-120 mm

## Slide 33: M041
EP treatment
  Bulk: 160 mm
Defect appears after EP on Eq1 w.r.t. coupler at 240° clockwise in beam direction
Defect is not visible after EBW
After EBW
After EP

## Slide 34: M041
3D-Microscopy investigation
  Defects Dimensions
  Morphology investigation
Next step:
  XRF (as for M042)
  X-Ray?
x1000
Height x 400
Transverse section profile
Border section profile
150-170 mm

## Slide 35: M042 - XRF Spectrum
Nb
Kα1 = 16615.1 eV
Kα2 = 16521.0 eV
Kβ = 18622.5 eV
Ar
Kα1 = 2957.7 eV
Kα2 = 2955.63 eV
Kβ = 3190.5 eV
Nb
Lβ1 = 2257.4 eV
Lβ2 = 2367.0 eV
X-Ray Source Parameters | 
X-Ray Energy [keV] | 22.3
Current [mA] | 1
Collimator diameter [mm] | 3
X-Ray (next step)
  Investigation presence of light element by “Intensity” test
IR Thermography (now on-going)
  Investigation of the thermal response during heating/cooldown of the defect
XRF
  Technique sensitive to high Z materials (Z > 14)
  No clear indication of foreign materials
  Diffraction effects visible in the defect region
    Light element contribution?
    Niobium composite?
    Other?

## Slide 36: M042 - XRF Spectrum
Nb
Kα1 = 16615.1 eV
Kα2 = 16521.0 eV
Kβ = 18622.5 eV
Ar
Kα1 = 2957.7 eV
Kα2 = 2955.63 eV
Kβ = 3190.5 eV
Nb
Lβ1 = 2257.4 eV
Lβ2 = 2367.0 eV
X-Ray Source Parameters | 
X-Ray Energy [keV] | 22.3
Current [mA] | 1
Collimator diameter [mm] | 3
X-Ray (next step)
  Investigation presence of light element by “Intensity” test
IR Thermography (now on-going)
  Investigation of the thermal response during heating/cooldown of the defect
XRF
  Technique sensitive to high Z materials (Z > 14)
  No clear indication of foreign materials
  Diffraction effects visible in the defect region
    Light element contribution?
    Niobium composite?
    Other?

## Slide 37: M042 - XRF Spectrum
Nb
Kα1 = 16615.1 eV
Kα2 = 16521.0 eV
Kβ = 18622.5 eV
Ar
Kα1 = 2957.7 eV
Kα2 = 2955.63 eV
Kβ = 3190.5 eV
Nb
Lβ1 = 2257.4 eV
Lβ2 = 2367.0 eV
X-Ray Source Parameters | 
X-Ray Energy [keV] | 22.3
Current [mA] | 1
Collimator diameter [mm] | 3
X-Ray (next step)
  Investigation presence of light element by “Intensity” test
IR Thermography (now on-going)
  Investigation of the thermal response during heating/cooldown of the defect
XRF
  Technique sensitive to high Z materials (Z > 14)
  No clear indication of foreign materials
  Diffraction effects visible in the defect region
    Light element contribution?
    Niobium composite?
    Other?

## Slide 38: Summary
The EP results on naked cavities are very encouraging
  M037 is near to specifications from 9.1 MV/m to 15.7 MV/m (> 70 %)
  M017 improved from 7.4 MV/m to 11.5 MV/m ( > 50 %)
    We are considering a further 50 mm EP
In our experience, EP etching rate is less reproducible w.r.t. BCP and we also observed cavity deformation and FF variation
Re-test of M028 will be done as soon as possible to understand the feasibility of rotating BCP for recovering jacketed cavities
The first new cavity treated with EP cycle largely overcome the ESS goal (but we are waiting for the mechanical and FF checks!). The second will be tested next week.
The new cavities with surface defects are under investigations with different not invasive techniques. If no clear evidence of foreign material will be detected, both will be grinded, treated with 20mm BCP + 30mm EP and then move towards VT test naked

## Slide 39
Thank you!